Confidential
Doc. No. FM-000 Rev. 1, Date: 10/31/2002
Status: Draft 

	ECMOre
	B62 Benedum Hall

Pittsburgh PA 15262

412-555-9873
	Standard Controlled Document Form

	
	
	Doc. No. FMEA-002
	Rev. 2

	
	
	Date: 04/16/04
	Status: Final



	Number: 
	FMEA-002
	Revision:
	2

	Title:
	Failure Modes and Effects Analysis

	UNCONTROLLED DOCUMENT UNLESS ISSUED WITH A RED STAMP

This document is the confidential property of ECMOre and may not be reproduced without prior written consent.

	Revision Approvals:
	Date

	Originator
	Desiree Bonadonna
	04/16/2004

	[Originating dept. approval]
	Bioengineering
	04/16/2004

	[Other dept. approval]
	Electrical Engineering
	04/16/2004

	Quality Assurance
	Laura Gilmour
	04/16/2004

	General Management
	Apryle Craig
	04/16/2004


1.0 Purpose

1.1. Organizes potential failure modes and effects of those failures

2.0 Scope

2.1. Effects hazard analysis

3.0 References

3.1 FMT-001

3.2 FMEA-001
4.0 Failure Mode and Effects Analysis
5.0 Revision History

5.1. FMEA-001; Revision 1; 12/04/03
5.2. FMEA-002; Revision 2; 04/16/04  FINAL
	function or component
	failure mode
	effect on system
	possible hazards
	risk index
	user detection means
	applicable control

	Transducer Line
	Extreme Temperatures
	Crack
	Minor Leak
	17
	Fluid dripping from connector
	Maintain reasonable room temperature

	 
	Extreme Temperatures
	Crack
	Break Sterility
	15
	Positive blood cultures
	 

	 
	Extreme Temperatures
	Crack
	Pressure Drop
	17
	Pressure Alarm
	Systematic checking of circuit

	 
	Extreme Temperatures
	Crack
	Disconnect-Major Leak
	12
	Pressure Alarm, fluid
	 

	 
	Extreme Temperatures
	Expansion
	Disconnect-Major Leak
	12
	Pressure Alarm, fluid
	Connector material able to withstand wide temperature ranges

	 
	Extreme Temperatures
	Melt
	No pressure output to monitor
	17
	Pressure Alarm
	 

	 
	Extreme Temperatures
	Melt
	Interruption of circuit flow
	12
	Pressure Alarm
	 

	 
	Incompatibility of material with blood
	blood reaction
	Patient Allergy
	10
	Adverse Patient Reaction
	Research compatibility of connector material before purchase

	 
	Incompatibility of material with blood
	blood reaction
	Rejection
	8
	Positive blood cultures
	Test patient for material allergies

	 
	Incompatibility of material with blood
	blood reaction
	Blood Clotting
	8
	Pressure alarm
	Maintain proper ACT's

	 
	Incompatibility with circuit
	unreliable connection
	Disconnect-Major Leak
	10
	Pressure Alarm, fluid
	research current circuits and ensure universal connections

	 
	Incompatibility with circuit
	Product useless
	No pressure readings
	11
	Unable to fit connector to tubing
	 

	Function or Component
	Failure Mode
	Effect on System
	Possible Hazards
	Risk Index
	User Detection Means
	Applicable Control

	Pressure Line Tubing
	Extreme Temperatures
	Crack
	Pressure Drop
	17
	Pressure Alarm
	Maintain reasonable room temperature

	 
	Extreme Temperatures
	Crack
	No pressure output to monitor
	17
	Pressure Alarm
	Systematic checking of circuit

	 
	Extreme Temperatures
	Expansion
	Interruption of pressure reading
	12
	Pressure Alarm
	Tubing material able to withstand wide temperature ranges

	 
	Extreme Temperatures
	Melt
	Interruption of pressure reading
	12
	Pressure Alarm
	 

	 
	Incompatibility
	unreliable connection
	Disconnect
	18
	Pressure Alarm, pressure line fluid
	research current circuits connections and pressure transducers

	 
	Incompatibility
	Product useless
	No pressure readings
	11
	Unable to fit tubing to connector or into pressure transducer
	 and ensure universal connections

	 
	Kink
	Increase pressure line pressure
	inaccurate reading
	9
	Pressure Alarm
	Maintain reasonable tubing length

	Transducer
	puncture
	hole in diaphragm 
	blood loss
	12
	blood seen in tubing connecting transducer to pressure monitor
	choose transducer approved for use in ECMO circuit

	 
	electrical malfunction
	incorrect voltage 
	inaccurate reading
	12
	unable to calibrate
	calibration/set-up examination

	Function or Component
	Failure Mode
	Effect on System
	Possible Hazards
	Risk Index
	User Detection Means
	Applicable Control

	Audible Alarm
	short in wires
	No audible alarm
	No attention drawn to pressure deviations
	15
	Abnormal read-out without audible alarm
	strain relief for wires

	 
	rusted wiring
	No audible alarm
	No attention drawn to pressure deviations
	17
	Abnormal read-out without audible alarm
	insulated wiring

	
	magnetic field presence
	No audible alarm
	No attention drawn to pressure deviations
	17
	Abnormal read-out without audible alarm
	protective enclosure

	Digital Pressure Display
	short in wires
	No digital display
	unable to assess patient's wellbeing
	15
	Abnormal read-out without audible alarm
	strain relief for wires

	 
	rusted wiring
	No digital display
	unable to assess patient's wellbeing
	17
	Abnormal read-out without audible alarm
	insulated wiring

	 
	magnetic field presence
	No digital display
	unable to assess patient's wellbeing
	17
	Abnormal read-out without audible alarm
	protective enclosure

	 
	short in wires
	incorrect readout
	possible misdiagnosis
	15
	Abnormal read-out without audible alarm
	Visual and audible alarm independent of display

	 
	rusted wiring
	incorrect readout
	possible misdiagnosis
	15
	Abnormal read-out without audible alarm
	 

	 
	magnetic field presence
	incorrect readout
	possible misdiagnosis
	15
	Abnormal read-out without audible alarm
	 

	Visual Alarm
	electrical malfunction
	No visual alarm
	unable to discern which pressure is out of physiological range
	15
	Audible alarm but no visual distinction
	Visual and audible alarm independent of display

	Function or Component
	Failure Mode
	Effect on System
	Possible Hazards
	Risk Index
	User Detection Means
	Applicable Control

	 
	short in wires
	No visual alarm
	unable to discern which pressure is out of physiological range
	15
	Audible alarm but no visual distinction
	strain relief for wires

	 
	rusted wiring
	No visual alarm
	unable to discern which pressure is out of physiological range
	15
	Audible alarm but no visual distinction
	insulated wiring

	
	magnetic field presence
	No visual alarm
	unable to discern which pressure is out of physiological range
	15
	Audible alarm but no visual distinction
	protective enclosure

	Programmable Range
	electrical malfunction
	does not save programmed range
	will not visually or audibly alarm when outside of range
	15
	display shows value outside of physiological range but no alarm goes off
	alarm independent of display

	
	short in wires
	does not save programmed range
	will not visually or audibly alarm when outside of range
	15
	display shows value outside of physiological range but no alarm goes off
	strain relief for wires

	 
	rusted wiring
	does not save programmed range
	will not visually or audibly alarm when outside of range
	15
	display shows value outside of physiological range but no alarm goes off
	insulated wiring

	 
	magnetic field presence
	does not save programmed range
	will not visually or audibly alarm when outside of range
	15
	display shows value outside of physiological range but no alarm goes off
	protective enclosure

	 
	 
	 
	 
	 
	 
	 


FORM: F0005 Standard Controlled Document Form.doc 

Page 1 of 5
FMEA-002


Page 5 of 5

