Confidential
Doc. No. FM-000 Rev. 1, Date: 10/31/2002
Status: Draft 

	PediaCool, Inc.
	749 Benedum Engineering

University of Pittsburgh

Pittsburgh, PA 15261
	Standard Controlled Document Form

	
	
	Doc. No. FTA0001-0001
	Rev. 1

	
	
	Date: 4/17/04
	Status: Draft



	Number:
	FTA 0001-0001
	Revision:
	1

	Title:
	Fault Tree Analysis for a Neonatal Heat Exchanger

	UNCONTROLLED DOCUMENT UNLESS ISSUED WITH A RED STAMP

This document is the confidential property of PediaCool, Inc. and may not be reproduced without prior written consent.

	Revision Approvals:
	Date

	Originator
	Kimberly Albrecht
	4/17/04

	[Originating dept. approval]
	Erin Aghamehdi
	4/17/04

	[Other dept. approval]
	Adam Abdulally
	4/17/04

	Quality Assurance
	Rebecca Hrutkay 
	4/19/04

	General Management
	PediaCool, Inc.
	


1.0 Purpose

1.1. The purpose of this document is to determine the possible routes of failure of a neonatal heat exchanger.

2.0 Scope

2.1. Re-design of a heat exchanger to rapidly induce hypothermic conditions.

3.0 References

3.1. N/A

4.0 Revision History

4.1. Revision 1 – FTA 0001-0001 – 4/17/04 - Polypropylene fiber bundle design.

[image: image1]
Rewarming of the patient








Kink in water side tubing





Crack in the potting material





Bad seal between shell and potting material





Bad seal between cap and shell





Water side tubing connector cracked or broken





PVC shell cracked or broken





Water side tubing improperly attached





Failure of health care worker to recognize clotting formation





Failure of a pressure sensor





Rupture of microfibers





Crack/broken cap





Blood side tubing connected improperly











Patient Injury/Death





Clotting Complications





Patient Blood Loss and Blood Contamination











FORM: F0005 Standard Controlled Document Form.doc 

Page 1 of 1
Document 1




Page 1 of 2

