Failure Mode Effects Analysis

Volumetric Airflow Gauge

Team Members:

Matthew Chakan

Justin Kiswardy

Michael Nilo

Principal Investigator:

Guy Guimond

Submittal Date:

April 17, 2007

Table 1. Risk Classification, Based on Occurrence Probability and Severity of Consequences

|  | Consequences |  |  |  |
| --- | --- | --- | --- | --- |
| Occurrence | Negligible | Marginal | Critical | Catastrophic |
| Frequent | A | C | D | D |
| Occasional | A | B | D | D |
| Remote | A | A | B | C |
| Improbable | A | A | B | B |
| Incredible | A | A | A | A |

Table 2. Interpretation of Risk Level

| Risk Level | Interpretation |
| --- | --- |
| A | Negligible risk |
| B | Tolerable risk |
| C | Undesirable risk |
| D | Intolerable risk |

Table 3. Failure Mode Effects Analysis

| Function or Component | Failure Mode | Effect on System | Possible Hazards | Risk Index | User Detection Means | Applicable Controls |
| --- | --- | --- | --- | --- | --- | --- |
| Electrical circuitry | Incomplete circuit, resistor/thermistor malfunction | Inaccurate volume display | Over/under-inflation by user as result of incorrect volume display | B | Visual inspection; realization of over/under-estimation | Suggest periodic maintenance of the VAG; use electrical components w/broad temperature ranges |
| LED malfunction | Burn out or incomplete circuit | Loss of power indicator/LED metronome | User doesn’t know when to replace power source; inaccurate timing of inflation/deflation cycles | B | Realization of loss of power/inaccurate timing | Suggest time-frame for battery life & periodic maintenance; incorporate LEDs w/low voltage demands |
| Casing | Degradation due to repetitive use/exposure to extreme enviro. | Exposes circuitry/loose parts; electrical malfunction | Inaccurate volume display interpreted by the resuscitator; microbial buildup inside device | C | Visual inspection | Strong-lightweight material, recommend disposal after 6000 cycle use |
| Reset trigger | Malfunction; no/incorrect reset after each cycle | Improper volume display to user | Over/under-inflation of patient’s airway | C | User awareness of airflow levels | Use microprocessor which minimizes error in reset function |
| LED display | Malfunction/ improper calibration | Incorrect volume display to user | Over/under-inflation of patient’s airway | C | Visual inspection, user awareness of airflow | Recommend periodic maintenance; QC calibration test of each device |
| Tubing | Crack/leak in tubing due to physical/enviro. damage | Compromised airflow to patient | Delayed/failure to resuscitate patient | D | Visual/ auditory inspection | Set acceptable temp. range for device use; incorporate tubing w/broad temp. range; suggest periodic maintenance |
| Inlet/outlet ports | Inadequate seal between device and BVM | Airflow leak resulting in inaccurate volume display | Under-inflation of patient’s airway; delayed time to resuscitation | C | Realization of leak due to visual/auditory inspection | Create device compliant with standard BVM tubing dimensions; QC test for portal diameters |
| Function or Component | Failure Mode | Effect on System | Possible Hazards | Risk Index | User Detection Means | Applicable Controls |
| Casing | Rubber gasket/seal failure; water/liquid enters system | Damaged electrical circuitry; compromised functionality of volumetric display | Inaccurate volume display interpreted by resuscitator; removal of device | B | Functionality of device; visual inspection | Minimize surface area of gaskets used; incorporate backup adhesive seal |
| Battery Housing | Break/incomplete seal of face plate | Environmental damage to batteries; increased power loss to circuitry | Decreased battery life | A | Realization of decreased battery life; visual inspection | Manufacture of face plate that requires pin for attachment as well as backup catch mechanism |
| Electrical circuitry | Improper/loose fit between electrical circuitry and casing; attachment pin failure | Movement of electrical circuitry and LED display inside of casing | Damage to loose electrical components inside casing; difficult for user to read display | A | Visual inspection | Utilize reinforced pin system to secure electrical circuit to casing; QC test to verify dimensions of casing |