Initial Hazard Analysis

Volumetric Airflow Gauge

Team Members:

Matthew Chakan

Justin Kiswardy

Michael Nilo

Principal Investigator:

Guy Guimond

Revision Date:

April 17, 2007

Potential Hazards & Risk Assignment:

- Device components enter patients airway - Catastrophic

- Device impedes/blocks exhalatory air flow - Catastrophic

- Device restricts/minimizes air from bag to patients airway - Catastrophic

- Device underestimates tidal volume - Critical

- Device overestimates tidal volume - Critical

- Device doesn’t register/display airflow volume – Marginal

- Device improperly displays inflation/deflation cycle time - Marginal

- Patient contracts infection from device - Marginal

Causes of Potential Hazards:

- Mass flow device tubing fails/cracks allowing direct access of electrical components to airway.  Electrical components break from circuit due to trauma/damage to device.

- Device is connected on the wrong side (distal) of the one-way flow valve

- Tubing malfunction/break, improper seal formed between device and BVM tubing

- Improper calibration of device; leak in tubing or seal on proximal side of device; electrical malfunction

- Improper calibration of device; device does not reset completely to “zero” after each cycle; electrical malfunction

- Battery failure; electrical circuit break; LED display malfunction; no power supplied to circuitry

- The LED metronome is not calibrated properly; circuit break/ microcontroller failure to LED metronome

- Improper disinfection of device;  casing failure allowing microbes to enter circuitry of device

System Responses:

- The device stops operating properly; bag pressure may increase/decrease; failure to resuscitate patient; patient asphyxiated by components in airway.

- Increased pressure in patients lungs; lung-tissue damage; delayed resuscitation time; user may overcompensate for increased pressure observed, therefore under-inflate patients lungs.

- Device displays inaccurate volume to user; user continues to inflate based on device display and patient receives inadequate air supply; delayed or failure to resuscitate

- Underestimated volumetric reading; excessive tidal volumes delivered to patient resulting in lung tissue damage, gastric distension, etc.

- Overestimated volumetric reading; inadequate tidal volumes delivered to patient resulting in delayed time to resuscitate/failure to resuscitate.

- User may not rely on device leading to over/under inflation observed with standard BVM use.

- The user may become disoriented w/ timing of inflation/deflation cycle during rescue; delayed time to resuscitate; user may not rely on device.

- Patient develops immune response

Table 1. Identification of Risk Level Acceptance

| Risk Level | Acceptable Risk |
| --- | --- |
| Negligible | Completely Acceptable |
| Marginal | Acceptable |
| Critical | Undesired |
| Catastrophic | Completely Unacceptable |