Risk Management Summary

## Volumetric Airflow Gauge

Team Members:

Matthew Chakan

Justin Kiswardy

Michael Nilo

Principal Investigator:

Guy Guimond

Submittal Date:

April 17, 2007

Table 1. Risk Classification, Based on Occurrence Probability and Severity of Consequences

|  | Consequences |  |  |  |
| --- | --- | --- | --- | --- |
| Occurrence | Negligible | Marginal | Critical | Catastrophic |
| Frequent | A | C | D | D |
| Occasional | A | B | D | D |
| Remote | A | A | B | C |
| Improbable | A | A | B | B |
| Incredible | A | A | A | A |

Table 2. Interpretation of Risk Level

| Risk Level | Interpretation |
| --- | --- |
| A | Negligible risk |
| B | Tolerable risk |
| C | Undesirable risk |
| D | Intolerable risk |

Table 3. Risk Management Summary

| Hazard | Consequences | Cause | Probability of Harm | Risk | Method of Control | Control Measures | New Probability | New Risk |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Insufficient air flow to patient | Critical | Insufficient VAG training/improper attachment | Remote | D | Process | Provide adequate training and instructive materials | Improbable | D |
| Device components enter patient’s airway | Catastrophic | Tubing degradation/fracture | Remote | D | Design | Strong material chosen and adequate tolerances designed | Improbable | B |
| Over/under inflation of patient’s airway | Critical | Device displays incorrect volume | Occasional | C | Design | Periodic maintenance of VAG; use electrical components w/broad temp. range | Remote | B |
| Inaccurate volume display | Critical | Rubber gasket/seal failure; water/liquid enters system | Remote | C | Design | Minimize gasket surface area used; incorporate backup adhesive seal | Remote | B |
| Over/under inflation of patient’s airway | Critical | No/incorrect reset after each cycle | Occasional | C | Design | Use of microprocessor to minimize reset function error | Improbable | B |
| Insufficient airflow to patient | Critical | Crack/leak in tubing | Remote | D | Design | Incorporate tubing w/broad temp. range; periodic maintenance | Improbable | B |
| Under-inflation of patient’s airway | Critical | Inadequate seal between device and BVM | Remote | C | Design | Comply with standard BVM tubing dimensions; QC test for portal diameters | Improbable | B |
| Patient develops immune response | Marginal | Improper device disinfection | Remote | B | Process | Allow sufficient disinfection time in alcohol | Improbable | A |
| Hazard | Consequences | Cause | Probability of Harm | Risk | Method of Control | Control Measures | New Probability | New Risk |
| Improper inflation/def. Cycle | Marginal | LED metronome malfunction | Remote | B | Design | QC test on LED’s. recommend periodic maintenance | Improbable | B |
| Inoperable device | Marginal | Power LED indicator failure, user doesn’t replace batteries | Remote | A | Design | QC test on LED’s, recommend periodic maintenance | Improbable | A |
| Inaccurate volume display, over/under inflation | Critical | Thermstor malfunction | Occasional | C | Design | Thermistor must be protected by housing | Improbable | B |
| Inaccurate volume display | Critical | Improper calibration of mass flow device | Occasional | C | Design | QC test on device, exact methodology for calibration steps | Remote | B |