Anterior Cervical Plate Hazard Analysis

- Causes of hazardous conditions

  - Development

    - Errors

      - Incorrect labeling of plates and screws

      - Poorly designed screws could not be incorrectly sized for the plate and vice versa (poor tolerance stack-up)

      - Incompatibility (size, material) with surgical tools and accepted practices (incision size, retraction systems, screw drivers)

      - Rotating screws or translating plates (with respect to screw head placement) could inhibit proper fusion

      - Lateral wing screw sizing could be too long for anatomy

      - Lateral wing screw diameter could weaken the plate

    - Development Response

      - Ensure clarity of labels via physician training/feedback

      - Enforce tight tolerances

      - Recommend use of specific surgical equipment with plate

      - Recommend specific static or rotating screws

      - Manufacture lateral screws for specific use in the wings

  - Environment

    - Errors

      - Broach of sterile packaging would necessitate pre-implant sterilization

      - Formation of a galvanic cell with platinum, gold or graphite

      - Once removed from packaging, plates need to survive autoclaving

    - Response to Environment

      - Use long-lasting adhesive for sterile packaging application

      - Ensure that plate is autoclavable

      - Recommend cool storage ( <37°C) prior to implantation

  - Implantation

    - Errors

      - Cervical musculature, spinal cord, nerve roots, vasculature and nearby organs (trachea, esophagus, epiglottis, thyroid glands, parathyroid glands) could be injured due to untrained or poorly executed surgical technique

      - Incorrect plate size choice (with respect to patient anatomy) due to improper labeling or poor sizing tolerance could result in instability or immediate post-operative failure of the construct

      - Incorrect plate-screw sizing could similarly result in acute or chronic instability

      - Improper length screw shafts could extend past the posterior cortex of a vertebral body into the spinal cord or nerve roots

    - Response to Implantation errors

      - Conduct workshops for surgeons illustrating potential hazards and proper techniques to minimize the risk

      - Clearly label plates.  Make bin with custom slots for each plate size to preclude the chance of misplacement

      - Screw color-coding or clear marking for and control on sizing

      - See (1.c.ii.3)

  - Post- surgery shifting

    - Errors

      - Anterior screws securing the plate could shift excessively due to subsidence (axially), torsion (laterally) or a combined load.  The screw displacement would alter segmental kinematics, change local loading and potentially contribute to instability.  Both static and fatigue loading must be considered.

        - Screws displaced through the endplates (axially) or through the graft/vertebral body wall (laterally) could cause loosening of the plate.  This could result in bone graft loosening or plate ejection into the surrounding tissue

        - Micromotion allowed by screws (permitted by angle variability or translation along the plate) could inhibit fusion

      - Screw selection and use

        - Translational load sharing- screw translate along a slot in the plate.  With multiple use of this feature over months/years it could become stuck causing discomfort on the patient

          - This could also cause stress on the disk or possible disk movement.

        - Any screw that does not perform a complete fusion and allows translation or rotation, giving more neck freedom, causes much more risk

      - Mechanical failure of the screws—weak screws could deform due to the large transverse load of the plate at the screw neck

      - Mechanical failure of the plate—unaccounted for torsional loading, axial loading or bending could cause buckling of the plate itself

      - Plate (fragment) ejection could cause dislodged metal in the surrounding tissue

      - Inserted bone could dislodge if plate buckled

    - Post-surgical Response

      - (a) (b) Revision surgery

      - (a)-(c) Risk of instability which may warrant revision surgery

      - – (6) Revision surgery