FAILURE MODES AND EFFECTS ANALYSIS

Retractable Tubing Device for Oxygen Therapy

JMAN: Jamie Haney, Mike Smithula, Andrew Vidokle, Nathan Sendgikoski

| Function or Component | Failure Mode | Effect on System | Possible Hazards | Risk Index | User Detection Means | Applicable Control(s) |
| --- | --- | --- | --- | --- | --- | --- |
| Oxygen concentration tubing | Tubing becomes tangled or kinked | Oxygen is restricted and therefore is not doing its desired task | Oxygen uptake would be restricted and lung function would be reduced | 10 | Patient condition worsens or flow gauge reads 0 | Track for tubing inside housing chamber |
| Oxygen concentration tubing | Tubing becomes damaged or holes emerge | Reduced pressure and oxygen flow | Patient could become ill due to lack of oxygen | 10 | Patient condition worsens or flow gauge values are below normal | Try to reduce friction by not allowing tubing to rub together |
| Constant force spring for retraction | Spring is too stiff | Causes Tube to retract to fast | Tubing is removed from connection ports | 7 | Possible injury or discomfort to patient, detachment from oxygen source | Test springs before implementing them into the design |
| Constant force spring for retraction | Spring wears down and malfunctions | Oxygen Therapy tube will not retract anymore and is stuck in one position | Excess tubing is prone to tangling | 2 | User notices there is excess tubing exposed and tries to retract it but it will not retract | Set guidelines for users to know the maximum amount of use for the system |
| Locking mechanism | Plastic mechanism breaks | Oxygen tubing will not lock into one place and cannot extend | If mechanism is already locked then a removal of the tubing possible, Limits patient mobility | 11 | User notices that system will not lock into place | Use a type of plastic that with minimum deformability |
| Housing chamber | Housing chamber breaks | Pieces of system may obstruct or puncture the tubing; oxygen flow is compromised | Patient could become ill because of lack of oxygen | 1 | User can see if the housing is broken | Use a type of plastic that with minimum deformability |